Grating Testing Apparatus (GTA)

User's Manual

Casey Deen

8 February 2011

Transmission Testing

1. Prepare the apparatus for the test. Be sure to that the test sample is OUT of the beam. Make sure that all instruments are powered on. This includes:
  1. NI Motor Controller (MID-7604) - Ensure the ENABLE switch is ON
  2. Spectral Products DK480 Monochrometer
  3. Light source (either optical or IR source)
  4. SRS510 Lock-in Amplifier
  5. SR540 Chopper Controller
  6. Detector electronics (power supply and cooler if using IR detector)
2. Start the Labview motor control VI (GT_server_v2.vi). This VI waits for approximately 1 minute to establish a IP connection to the python program. After the connection is established, it will wait indefinitely for commands from the python program over the IP connection.
3. Open the transmission.py python script in the python IDLE. Press F5 to start.
  1. The script will load in the defaults from the grating_test.config.txt file.
  2. The motor controller, if configured correctly, will send the motor searching for its home position. Be sure to start the program with the detector arm clock-wise from its home location.
    1. If you fear the detector arm is on the wrong side of the “home” sensor, disable the NI MID-7604 motion controller using the “ENABLE” switch on the front panel. Then, use the knob on the front of the rotation stage to move the detector arm back on the correct side of the home sensor. The home sensor is located ~15 degrees away from the straight-through beam (in a counter-clockwise orientation when looking down the rotation axis from above.)
4. Once the rotation stage has found its home location, the program will ask you for the following information
  1. the name of the data file in which you wish to store output data.
  2. Wavelength starting and stopping points (in nm)
  3. Number of intermediate wavelength points to test
  4. Name of sample
  5. Email address (for notification)
  6. Desired Signal-to-Noise ratio
  7. Desired filter wheel location (see Appendix A)
  8. Monochromator slit width (in microns)
5. After gathering the required information, the program will do the follwing:
  1. Tune the monochromator to the beginning wavelength
  2. Open the shutter (move filter wheel to specified location)
  3. Move the detector arm through a small range of angles centered around the beam, in order to find the location of the maximum signal
  4. Close the shutter (move the filter wheel to blank position)
  5. Record the “background” level
  6. Open the shutter (move filter wheel to specified location)
  7. Record the “straight-through” reference measurement for each wavelength
  8. Halt for user input
    1. Place transmission sample in the beam
    2. Press “ENTER” to continue test
  9. Record a measurement for each wavelength of light transmitted through the material
    1. Calculate the absolute transmission factoring in the background and reference measurement
  10. Save raw data, transmission curve, and raw data plot

Appendix A: Filter wheel Assignments
Filter Wheel Number	Contents	Wavelength Range
1	Empty	0->inf
2	Melles Griot Filter	< 1.9 microns
3	Empty	0->inf
4	AB3190	< 3190 nm
5	AB330	< 3300 nm
6	Blank	none